____________________________

# SPECIFICATIONS

# Vertical Measurement Ranges

Autoranging up to: +160KA or +160um

____________________________

# Horizontal Measurement Ranges and Sampling

## Max. Scan   	Samples Per Micrometer

## Length	     8-sec. Scan  40-sec. Scan  Scan Direction

10,000um	 --	      0.2		right only

2,000um	0.2	      1.0		left or right

400um	1.0	      5.0		left or right

80um	5.0	    25.0		left or right

Scan direction (right or left) and shorter scan length can be programmed by the operator.

Maximum scan length: 10,000um (10mm) available in

micrometer mode only.

____________________________

# Scan Menu Parameters (Nine Menus)

Scan length

Sampling rate

Scan direction

Stage Position (Programmable X-Y Stage only)

____________________________

Resolution

Vertical:		5A with + 160KA range

5nm with + 160um range

Horizontal:	400A

____________________________

Digital Conversion Rate

50 samples/second

____________________________

Stylus Radius

Standard:	12.5um or 5.0um

Optional:	1.5um to 2.5um, and submicron

____________________________

Stylus Tracking Force

Adjustable between 1 and 25mg and continuously monitored and displayed on CRT. Forces above 25mg may also be set.

____________________________

Scan Method

Stylus moves, stage remains stationary

____________________________

Sample Leveling

Automatically computed. Manual coarse-leveling adjustment.

____________________________

# Precision X-Y Stage

Maximum Sample Dimensions

Thickness: with theta table: 17mm (0.67”)

without theta table: 21mm (0.82”)

Diameter:   Measures to center of 162mm (6.4”) sample

Sample Stage Movement

X-Axis		100mm		(4.0”)

Y-Axis		81mm		(3.2”)

Z-Axis		21mm		(0.82”)

360-degree theta rotation

Sample Location

Precision manual knobs for X-Axis and Y-Axis stage movement; servocontrol up/down buttons for Z-Axis stage movement.

____________________________

# PROGRAMMABLE X-Y STAGE

Operating Modes

Manual

Semi-Automatic

Automatic

Maximum Number of Programmable Locations:

9 + “home” (0)

Locator Block Positions:

3”, 100mm, 125mm, 150mm

Maximum Sample Dimensions:

Thickness: with theta table: 10mm (0.4”)

without theta table: 17mm (0.67”)

Diameter:   Measures to center of 162mm (6.4”) sample

Stage Movement Range

X-Axis			99mm (3.9”)

Y-Axis			81mm (3.2”)

Z-Axis			17mm (0.67”)

180-degree theta rotation

Stage Movement Rate

Fast rate		13mm/sec (0.5”/sec)

Slow rate		0.13mm/sec (0.005”/sec)

Minimum increment:	1.6 microns (using motorized

control)

Positional Repeatability: +25 microns each axis

Vacuum Requirements: 100 torr, minimum